Magyar Nemzeti Bank

Központi beszerzés

AJÁNLATI DOKUMENTÁCIÓ

„Teljeskörű mérlegelési technika kialakítása”

KBE/67/2010

tárgyú beszerzési eljárásban

Részei:

1. melléklet: Feladatleírás

2. melléklet: Vállalkozási szerződés

3. melléklet: Ártáblázat (Felolvasólap, Tételes árlebontás)

4. melléklet: Ajánlattételi nyilatkozatminta

5. melléklet: Nyilatkozat felelősségbiztosításról

AJÁNLATI DOKUMENTÁCIÓ 1. szÁMÚ melléklet

FELADATLEÍRÁS

Komplett mérlegelési technika kialakítása
a bankjegykötegek ellenőrzésére, a tizedelés kiváltására

A Magyar Nemzeti Bank Logisztikai Központjában működő bankjegyfeldolgozó rendszerek a feldolgozott bankjegyekből címletenként kötegeket alakítanak, melyek jelenleg tizedeléses kézi ellenőrzésre kerülnek az esetleges csomag eltérések kiszűrésére.

(Tizedelés: a fóliázott bankjegykötegben lévő bankjegycsomagok /1 bankjegycsomag = 100 db bankjegy/ számának ellenőrzése a bankjegycsomagokon található bankjegyszalagok számbavétele alapján.)

Az ellenőrzés fontos része a feldolgozási folyamatnak, mivel az esetleges eltérések komoly presztízs és anyagi kockázatot hordoznak az MNB számára.

A tizedeléses kézi ellenőrzési folyamat kiváltására alakíttatjuk ki a teljes körű mérlegelési rendszert.

A mérlegelés a feldolgozási folyamatba integrálva működne úgy, hogy a csomagolóból kijövő bankjegykötegek (váltakozva minden címletből érkezik) a szállítószalag végénél lévő mérlegen egyesével automatikusan súlymérésre kerülnek és beépített szoftver értékeli a mért súlyt. Tűréshatáron belüli eredmény esetén a köteg a gyűjtőasztalra kerül továbbításra, ellenkező esetben egy elkülönítőbe.

A törvényes forintbankjegy fizetőeszközök (6 címlet) azonos méretűek, iránymutatásul közöljük, hogy 1 köteg súlya ~1000 g. A kötegek súlyának pontos megállapítása és a tűréshatárok kialakítása, a pályázat nyertesének feladata az általunk biztosított címletenkénti max. 100 bankjegyköteg mérlegelésével (előre egyeztetett időpontban).

Az alábbi elvárásokat fogalmazzuk meg a mérlegelésre kialakítandó rendszerrel kapcsolatban:

- A bankjegyfeldolgozó rendszer csomagolójából kikerülő bankjegykötegek egyesével, (leggyorsabb alakítási sebesség 10 másodperc) automatikusan, emberi közreműködés nélkül mérlegelésre kerüljenek a csomagolóhoz illesztett mérlegelő rendszeren.

- A mért súly értékelésével a rendszer válogassa szét és különítse is el a kötegeket, megfelelő illetve tűréshatáron kívül eső kategóriák szerint. Súlyeltérés esetén fény és hang, figyelmeztető jelzés adása a helyszínen. A megfelelő súlyú kötegeket a berendezés automatikusan továbbítsa a kötegek gyűjtőhelyére.

- A kötegek (1000 db bankjegy) mérésének elvárt pontossága: legalább 50 db bankjegy eltérés detektálása. Az MNB-t érintő anyagi és presztízs veszteség elkerülése érdekében biztosítani és szavatolni kérjük  a rendszer teljes élettartamára vonatkozóan, hogy a megfelelő súlyú kötegek közé nem kerül tűréshatáron kívül eső köteg. A súlyhatáron kívül eső kötegek között az 1000 db bankjegyet tartalmazó kötegek minimálisan tolerálhatók (a nem megfelelő súly miatt elkülönített összes köteg számának 30 százalékáig tolerálhatónak tartjuk, ha azok között megfelelő, 1000 db bankjegyet tartalmazó kötegek előfordulnak.)

- A mérleg használati helyén meglévő normál működés melletti folyamatos rezonanciát a mérlegelési rendszernek ki kell küszöbölni, a rezgés nem befolyásolhatja a mérési eredményt.

- A mérési adatoknak megfelelő legpontosabb beállítások és későbbi (utólagos) finomítási beállítások biztosítása a szállító cég által a folyamatos mérési tapasztalatok alapján a vállalt garanciális időn belül.

- Minden mért adat fájlba tárolása az e célra kialakított adattárhelyen. Az adatok kinyerésének biztosítása USB porton keresztül, Microsoft Office 2007 által támogatott formában. Az adatok strukturáltan lekérdezhetők legyenek napi, heti, havi bontásban. Jelenítse meg napi szinten, illetve a tűréshatárok változtatásakor, hogy a mért súly értékelése milyen tűrési tartomány alapján történik. A fájlban tárolandó adatok:

- a mérés dátuma

- a mérés időpontja (óra, perc, másodperc)

- a mért súly gramm pontossággal

- jelzés arra vonatkozóan, ha a köteg a megfelelő súlyúaktól elkülönítésre került (a tényleges fizikai elkülönítés megjelenítése)

- A két géptermi helyiségben külön-külön meg kell valósítani a mérlegelési rendszert, tehát összesen 2 szerkezet felépítése szükséges.

- Naponta mérendő mennyiség géptermenként maximum 1500 bankjegyköteg, heti munkanapok száma: 4.

Konzultáció keretében biztosítjuk a kialakítás helyének, a működés körülményeinek megtekintését és a szükséges méretadatok (hely, mérleg magasság stb.) felmérését, továbbá címletenként 10 bankjegyköteg mérlegelését. A méréseket az ajánlattevők saját eszközeikkel végzik.

AJÁNLATI DOKUMENTÁCIÓ 2. SZÁMÚ MELLÉKLET

VÁLLALKOZÁSI SZERZŐDÉS

TERVEZET

AJÁNLATI DOKUMENTÁCIÓ 3. SZ. MELLÉKLET

ÁRTÁBLÁZAT(Felolvasólap, Tételes árlebontás, Referenciamunkák)

AJÁNLATI DOKUMENTÁCIÓ 4. SZ. MELLÉKLET

ajánlattételi Nyilatkozat

Alulírott ……………...……….………, mint a(z) ……………………………..… (cégnév, székhely) Ajánlattevő cégjegyzésre jogosult képviselője felelősségem teljes tudatában az alábbiakról

nyilatkozom,

        - Az ajánlatkérésben és az ajánlati dokumentáció 2 sz. melléklete szerinti Vállalkozási szerződés tervezetben foglalt valamennyi formai és tartalmi követelményt megismertem és változatlan tartalommal elfogadom;

      - Az ajánlatunkban bemutatott eszközök megfelelnek az Ajánlati dokumentáció részét képező Feladatleírásban meghatározott követelményeknek.

3.	Kötelezettséget vállalok arra, hogy az eljárás nyerteseként az Ajánlatkérővel megkötjük az Ajánlati dokumentáció 2. sz. mellékleteként csatolt Vállalkozási szerződést, és a megrendelt eszközök beüzemelését az ajánlatunkban rögzített áron, a kért teljesítési határidőben elvégezzük.

4.	Az általam jegyzett cégre nem állnak fenn az ajánlatkérés 7. pontjában megfogalmazott kizáró okok.

5.	Felelősségem tudatában nyilatkozom, hogy a szerződés teljesítésébe nem vonok be olyan alvállalkozót aki/amely az ajánlatkérés 7. pontjában meghatározott kizáró okok hatálya alatt áll.

6.	A benyújtott ajánlatomat az ajánlattételi határidőt követő 60 napig fenntartom;

Jelen nyilatkozatot a Magyar Nemzeti Bank, mint ajánlatkérő által a „Teljeskörű mérlegelési technika kialakítása KBE/067/2010” tárgyában indított ajánlatkérésban az ajánlat részeként teszem.

Budapest, 2010. …………………………

….………………………………..

(Ajánlattevő cégszerű aláírása)

AJÁNLATI DOKUMENTÁCIÓ 5. SZ. MELLÉKLET

NYILATKOZAT FELELŐSSÉGBIZTOSÍTÁSRÓL

„A” típusú nyilatkozat

Alulírott ………………….…………., mint a(z) ………………...………….……………… (cégnév, székhely) Ajánlattevő cégjegyzésre jogosult képviselője ezennel felelősségem tudatában

n y i l a t k o z o m,

hogy a Magyar Nemzeti Bank, mint Ajánlatkérő által indított „Teljeskörű mérlegelési technika kialakítása KBE/067/2010” tárgyú egyéb beszerzési eljárás nyertes Ajánlattevőjeként legkésőbb a Vállalkozási szerződés aláírásától számított 10 napon belül a Vállalkozási szerződés szerinti tevékenységre - a harmadik személynek okozott károk fedezetét is magába foglaló – legalább 3.000.000,- Ft összegű felelősségbiztosítást kötünk, és azt az Ajánlatkérő részére bemutatjuk. Kötelezettséget vállalunk arra, hogy a jelen nyilatkozat szerinti felelősségbiztosítást a szerződés időbeli hatálya alatt legalább a fenti összegben mindvégig fenntartom.

Kelt ………………., 2010……………….

……………………………..

(cégszerű aláírás)

- - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - -  - - - - -

„B” típusú nyilatkozat

Alulírott ………………….…………., mint a(z) ………………...………….……………… (cégnév, székhely) Ajánlattevő (közös ajánlattevő) cégjegyzésre jogosult képviselője ezennel felelősségem tudatában

n y i l a t k o z o m,

a Magyar Nemzeti Bank, mint ajánlatkérő által indított „Teljeskörű mérlegelési technika kialakítása KBE/067/2010” tárgyú egyéb beszerzési eljárásban, hogy Cégünk az ajánlatkérésben előírtaknak megfelelően legalább 3.000.000,- Ft összegű felelősségbiztosítással rendelkezik, mely a Megrendelő részére teljes körű felelősségbiztosítási fedezetet jelent a harmadik személynek okozott károkért, és ezt a Vállalkozási szerződés időbeli hatálya alatt mindvégig fenntartjuk.

Kelt ………………., 2010……………….

……………………………..

(cégszerű aláírás)

* A nem kívánt nyilatkozatot áthúzással, a kívánt nyilatkozatrészt cégszerű aláírással kérjük benyújtani.

Magyar Nemzeti Bank

Central Procurement

TENDER DOCUMENTATION

‘Implementation of complete weighing technique’

in procurement procedure

KBE/67/2010

It includes:

Annex 1: Task description

Annex 2: Contract for work done and materials supplied

Annex 3: Price schedule (Project fiche, Itemised price breakdown)

Annex 4: Sample of tenderer’s declaration

Annex 5: Declaration on liability insurance

ANNEX 1 TO TENDER DOCUMENTATION

TASK DESCRIPTION

Implementation of complete weighing technique
to check bundles of banknotes, to replace decimation

The banknote processing systems operating in the Logistics Centre of the Magyar Nemzeti Bank make bundles of the processed banknotes by denominations. At present, the bundles are checked by manual decimation in order to eliminate deviating packages, if any.

(Decimation means the checking of the number of the banknote packages /1 package of banknotes = 100 ea banknotes/ in the foil-wrapped bundle of money on the basis of taking stock of the banknote wrappers on the banknote packages.)

Checking is an important part of the processing procedure, because possible deviations carry serious reputation and material risks for the MNB.

The complete weighing system is to be developed in order to replace the manual decimation checking procedure.

Weighing would be integrated in the course of processing in a way that the banknote bundles leaving the wrapping unit (all denominations arrive alternately) are automatically weighed one by one on the scales at the end of the conveyor belt, and the installed software evaluates the weight measured. If the result is within the limit of tolerance, the bundle goes to the collecting table for forwarding, otherwise it is separated.

The forint banknotes that are legal tender (6 denominations) have identical dimensions. For your information we indicate that the weight of 1 bundle is ~1000 g. Establishing the exact weight of the bundles and determining the limits of tolerance shall be the task of the winner of the tender by weighing max. 100 bundles of banknotes of each denomination provided by us (at a pre-arranged date).

The requirements regarding the weighing system to be developed shall be as follows:

- The banknote bundles leaving the wrapping unit of the banknote processing system shall be weighed one by one (fastest forming speed: 10 seconds), automatically, without any human assistance, on the weighing system attached to the wrapping unit.

- By evaluating the measured weight, the system shall sort and separate the bundles, according to categories ‘suitable’ and ‘outside the limit of tolerance’. In the event of a weight deviation, a warning light and audible signal shall be given on the spot. The bundles with adequate weight shall automatically be forwarded by the machine to the collection place of the bundles.

- The required accuracy of weighing the bundles (1000 ea banknotes): detecting a deviation of at least 50 ea banknotes. In order to avoid material losses and loss of prestige of the MNB, it shall be ensured and guaranteed for the entire life of the system that no bundle outside the limit of tolerance is mixed with the bundles of adequate weight. Among the bundles out of the limit of tolerance the ones containing 1000 ea banknotes can be tolerated to a minimum extent (up to 30 per cent of the number of all bundles separated because of inadequate weight it is tolerable if there are suitable bundles containing 1000 ea banknotes among them).

- The weighing system shall eliminate the continuous resonance during the normal operation at the place of using the scales; the resonance shall not influence the result of weighing.

- Ensuring the most accurate adjustments and later (subsequent) fine-tunings by the supplier in line with the measuring data, on the basis of the continuous weighing experiences, within the undertaken guarantee period.

- Storing all the measured data in a data store set up for this purpose. Ensuring the recovery of data through USB port, in a format supported by Microsoft Office 2007. The data shall be retrievable in a structured manner, in a daily, weekly and monthly breakdown. The range of tolerance based on which the measured weight is evaluated shall be displayed at a daily level and upon changing the limits of tolerance. The data to be stored in the file:

- the date of weighing,

- the time of weighing (hour, minute, second),

- the measured weight with gram accuracy,

- an indication if the bundle has been separated from those with an adequate weight (displaying the actual physical separation).

- The weighing system shall be implemented in both machine rooms separately, i.e. altogether 2 devices shall be installed.

- The daily quantity to be weighed in each machine room is not more than 1500 banknote bundles; the number of working days per week: 4.

Inspecting the place of installation and the circumstances of operation as well as measuring the necessary dimensions (place, height of the scales etc.) and also weighing 10 banknote bundles of each denomination will be made possible within the framework of consultation. Tenderers shall use their own means for the measurements.

ANNEX 2 TO TENDER DOCUMENTATION

CONTRACT FOR WORK DONE AND MATERIALS SUPPLIED

DRAFT

ANNEX 3 TO TENDER DOCUMENTATION

PRICE SCHEDULE (Project fiche, Itemised price breakdown, References)

ANNEX 4 TO TENDER DOCUMENTATION

TENDERER’S DECLARATION

I, the undersigned …………………………, as the person vested with the power of representation of ……………………………… (company name, registered office), the tenderer, being fully aware of my responsibility,

hereby declare that

        - I have become acquainted with all the form and content requirements set forth in the invitation to bid and in the draft contract for work done and materials supplied in Annex 2 to the tender documentation, and accept them with an unchanged content;

      - the appliances presented in our tender comply with the requirements specified in the Task description that is a part of the tender documentation.

3.	I assume an obligation that, as winners of the procedure, we will conclude the contract for work done and materials supplied attached as Annex 2 to the tender documentation with the contracting authority, and we will carry out the commissioning of the appliances ordered at the price fixed in our tender, by the requested performance deadline.

4.	The grounds for exclusion described under point 7 of the invitation to bid do not apply to the company represented by me.

5.	Being aware of my responsibility I declare that I will not involve any subcontractor in the performance of the contract that comes under the ruling of the grounds for exclusion set forth under point 7 of the invitation to bid.

6.	My tender submitted is binding for 60 days from the tender deadline.

This declaration is a part of the tender submitted for the invitation to bid ’Implementation of complete weighing technique KBE/067/2010’ announced by the Magyar Nemzeti Bank as contracting authority.

Budapest, …………………………. 2010

….………………………………..

(tenderer’s authorised signature)

ANNEX 5 TO TENDER DOCUMENTATION

DECLARATION ON LIABILITY INSURANCE

Declaration type ‘A’

I, the undersigned …………………………, as the person vested with the power of representation of ……………………………… (company name, registered office), the tenderer, being aware of my responsibility,

d e c l a r e

that as the winning tenderer of the other procurement procedure ‘Implementation of complete weighing technique KBE/067/2010’ launched by the Magyar Nemzeti Bank as contracting authority we will sign a liability insurance contract for the activity set forth in the contract for work done and materials supplied, with an insured amount of at least HUF 3,000,000 – covering damages caused to third persons as well – within 10 days from signing the contract for work done and materials supplied, and will present this insurance contract to the contracting authority. We also undertake to maintain the liability insurance mentioned in this declaration at least with the aforementioned amount until the contract is valid.

Dated: ……………….., ……………….. 2010

……………………………..

(authorised signature)

- - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - -  - - - - -

Declaration type ‘B’

I, the undersigned …………………………, as the person vested with the power of representation of ……………………………… (company name, registered office), the tenderer (joint tenderer), being aware of my responsibility, hereby

d e c l a r e

that, in conformity with the requirements set forth in the invitation to bid, our company has a liability insurance with an insured amount of at least HUF 3,000,000 in the other procurement procedure ‘Implementation of complete weighing technique KBE/067/2010’, which means a full liability insurance cover for the contracting authority for damages caused to third persons, and we shall maintain this insurance until the contract for work done and materials supplied remains valid.

Dated: ……………….., ……………….. 2010

……………………………..

(authorised signature)

* Please, cross out the declaration that does not apply, and submit the one that does apply with your authorised signature.